## Sheet1
Weights calibrated measuring a pure LaB6 standard, bulk, polished | [10 empty] | Weights unchanged
Area of all peaks is fitted | [10 empty] | Weights changed
A constant shift is maintained across all peaks
Sigma is fixed for all peaks
Had to separate the peak fitting in two groups, otherwise it would not fit
La |  | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10
Mz1 |  | 67.402 | 66.7 | 68.409 | 65.62 | 66.059 | 70.293 | 67.348 | 69.292 | 70.181 | 69.641
M5O3 |  | 0 | 0.315 | 0.139 | 0 | 0 | 0.01 | 0.002 | 0.002 | 0 | 0
M4O2 |  | 20.985 | 10.218 | 4.968 | 12.186 | 13.537 | 25.905 | 19.307 | 31.73 | 14.185 | 24.374
M3N1 |  | 20.781 | 33.755 | 35.631 | 27.996 | 30.614 | 18.911 | 23.728 | 11.877 | 29.714 | 18.554
Mz1 | Looked fine with NIST calibration
M2N1 | Looked fine with NIST calibration
Mg | Looked fine with NIST calibration
M2N4 | Looked fine with NIST calibration
Line | Original w | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10 | Mean | Stdev | Er (%)
Mz1 | 1 | [10 empty] | 1 | bar: Meas 1 to Meas 10
M5O3 | 0.000262 | 0 | 0.0047226 | 0.0020319 | 0 | 0 | 0.00014226 | 2.9697e-05 | 2.8863e-05 | 0 | 0 | 0.00069554 | 0.0015498 | 222.8172
M4O2 | 0.00259 | 0.31134 | 0.15319 | 0.072622 | 0.18571 | 0.20492 | 0.36853 | 0.28668 | 0.45792 | 0.20212 | 0.34999 | 0.2593 | 0.11586 | 44.6816
M3N1 | 0.484 | 0.30831 | 0.50607 | 0.52085 | 0.42664 | 0.46343 | 0.26903 | 0.35232 | 0.17141 | 0.42339 | 0.26642 | 0.37079 | 0.11573 | 31.2111
M5O3 + M4O2 + M3N1 | 0.48685 | 0.61966 | 0.66399 | 0.59551 | 0.61234 | 0.66836 | 0.6377 | 0.63902 | 0.62935 | 0.62551 | 0.61642 | 0.63079 | 0.022545 | 3.5742
M2N1
Mg
M2N4
Fits whose peak areas vary, but whose sum is almost cosntant:
So we use the original ratios to diustribute the sum among the peaks, giving:
M5O3 + M4O2 + M3N1 | M5O3 | Meas 1: M4O2; Meas 2: M3N1; Meas 3: Sum
Original ratio | 0.00053815 | Meas 1: 0.0053199; Meas 2: 0.99414
Resulting weights: | 0.00033946 | Meas 1: 0.0033557; Meas 2: 0.62709; Meas 3: 0.63079
|  | Without fixing the peaks above: | [8 empty] | Fixing the peaks: