## Sheet1
Weights calibrated measuring a pure Mo standard, bulk, polished | [10 empty] | Weights unchanged
Area of all peaks is fitted | [10 empty] | Weights changed
A constant shift is maintained across all peaks
Sigma is fixed for all peaks
We do not touch Lg2 and Lg3 because our background is not good enough there and because they seem already well calibrated
Mo |  | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10
La1 |  | 747.112 | 747.174 | 741.22 | 737.402 | 751.518 | 738.973 | 742.792 | 745.065 | 747.303 | 740.706
Ll |  | 34.282 | 33.434 | 33.465 | 32.97 | 34.36 | 34.237 | 33.058 | 33.889 | 33.109 | 32.906
Ln |  | 10.401 | 12.24 | 12.417 | 10.425 | 10.562 | 10.134 | 12.033 | 9.591 | 11.462 | 12.166
Lb3 |  | 64.665 | 52.911 | 54.582 | 35.928 | 39.558 | 58.066 | 23.534 | 36.584 | 48.331 | 57.365
Lb4 |  | 0.068 | 14.602 | 8.668 | 26.145 | 24.408 | 9.02 | 42.906 | 27.616 | 18.865 | 11.17
Lb3 + Lb4 |  | 64.733 | 67.513 | 63.25 | 62.073 | 63.966 | 67.086 | 66.44 | 64.2 | 67.196 | 68.535
Lb1 |  | 345.952 | 339.55 | 340.498 | 337.31 | 336.377 | 334.314 | 326.866 | 337.678 | 339.428 | 338.894
Lb2 |  | 0.062 | 11.229 | 9.593 | 11.088 | 12.12 | 5.918 | 12.694 | 11.787 | 0.001 | 0
Lb7 |  | 13.791 | 6.99 | 7.317 | 9.386 | 10.295 | 4.853 | 11.979 | 9.654 | 15.907 | 14.869
Lb2 + Lb7 |  | 13.853 | 18.219 | 16.91 | 20.474 | 22.415 | 10.771 | 24.673 | 21.441 | 15.908 | 14.869
Lg5 |  | 2.163 | 0 | 0 | 0.001 | 0 | 8.39 | 0 | 0.029 | 2.863 | 5.217
Lg11 |  | 1.693 | 0 | 1.314 | 0.107 | 0.235 | 0 | 0.403 | 0.11 | 0.032 | 0
L2N3 |  | 0 | 4.493 | 4.108 | 4.488 | 1.219 | 0.001 | 1.32 | 4.616 | 1.431 | 0
Lg5 + Lg11 + L2N3 |  | 3.856 | 4.493 | 5.422 | 4.596 | 1.454 | 8.391 | 1.723 | 4.755 | 4.326 | 5.217
Lg1 |  | 3.333 | 2.452 | 1.907 | 2.722 | 3.086 | 2.392 | 2.982 | 2.026 | 2.477 | 3.513
Lg8 |  | 3.333 | 2.452 | 1.907 | 2.722 | 3.086 | 2.392 | 2.982 | 2.026 | 2.477 | 3.513
Lg1 + Lg8 |  | 6.666 | 4.904 | 3.814 | 5.444 | 6.172 | 4.784 | 5.964 | 4.052 | 4.954 | 7.026
Lg2 |  | 2.321 | 2.484 | 1.342 | 1.551 | 2.349 | 0.411 | 1.479 | 2.288 | 4.1 | 3.65
Lg3 |  | 2.302 | 0.993 | 1.481 | 1.218 | 1.484 | 3.607 | 1.599 | 1.357 | 0 | 0
Lg2 + Lg3 |  | 4.623 | 3.477 | 2.823 | 2.769 | 3.833 | 4.018 | 3.078 | 3.645 | 4.1 | 3.65
Line | Original w | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10 | Mean | Stdev | Er (%)
La1 | 1 | [10 empty] | 1 | bar: Meas 1 to Meas 10
Ll | 0.0415 | 0.045886 | 0.044747 | 0.045149 | 0.044711 | 0.045721 | 0.046331 | 0.044505 | 0.045485 | 0.044305 | 0.044425 | 0.045126 | 0.00069748 | 1.5456
Ln | 0.0128 | 0.013922 | 0.016382 | 0.016752 | 0.014137 | 0.014054 | 0.013714 | 0.0162 | 0.012873 | 0.015338 | 0.016425 | 0.01498 | 0.001396 | 9.3195
Lb3 | 0.06299 | 0.086553 | 0.070815 | 0.073638 | 0.048722 | 0.052637 | 0.078577 | 0.031683 | 0.049102 | 0.064674 | 0.077446 | 0.063385 | 0.017214 | 27.1574
Lb4 | 0.03669 | 9.1017e-05 | 0.019543 | 0.011694 | 0.035456 | 0.032478 | 0.012206 | 0.057763 | 0.037065 | 0.025244 | 0.01508 | 0.024662 | 0.016575 | 67.2077
Lb3 + Lb4 | 0.09968 | 0.086644 | 0.090358 | 0.085332 | 0.084178 | 0.085116 | 0.090783 | 0.089446 | 0.086167 | 0.089918 | 0.092527 | 0.088047 | 0.0028826 | 3.2739
Lb1 | 0.32736 | 0.46305 | 0.45445 | 0.45938 | 0.45743 | 0.4476 | 0.4524 | 0.44005 | 0.45322 | 0.4542 | 0.45753 | 0.45393 | 0.0064463 | 1.4201
Lb2 | 0.04509 | 8.2986e-05 | 0.015029 | 0.012942 | 0.015037 | 0.016127 | 0.0080084 | 0.01709 | 0.01582 | 1.3381e-06 | 0 | 0.010014 | 0.0073278 | 73.1781
Lb7 | 0.0004 | 0.018459 | 0.0093553 | 0.0098716 | 0.012728 | 0.013699 | 0.0065672 | 0.016127 | 0.012957 | 0.021286 | 0.020074 | 0.014112 | 0.004841 | 34.303
Lb2 + Lb7 | 0.04549 | 0.018542 | 0.024384 | 0.022814 | 0.027765 | 0.029826 | 0.014576 | 0.033217 | 0.028777 | 0.021287 | 0.020074 | 0.024126 | 0.0057606 | 23.8771
Lg5 | 0.0022 | 0.0028951 | 0 | 0 | 1.3561e-06 | 0 | 0.011354 | 0 | 3.8923e-05 | 0.0038311 | 0.0070433 | 0.0025163 | 0.0039194 | 155.7566
Lg11 | 0.0001 | 0.0022661 | 0 | 0.0017728 | 0.0001451 | 0.0003127 | 0 | 0.00054255 | 0.00014764 | 4.2821e-05 | 0 | 0.00052296 | 0.00081511 | 155.8636
L2N3 | 0.0027 | 0 | 0.0060133 | 0.0055422 | 0.0060862 | 0.0016221 | 1.3532e-06 | 0.0017771 | 0.0061954 | 0.0019149 | 0 | 0.0029153 | 0.0027238 | 93.4316
Lg5 + Lg11 + L2N3 | 0.005 | 0.0051612 | 0.0060133 | 0.007315 | 0.0062327 | 0.0019348 | 0.011355 | 0.0023196 | 0.006382 | 0.0057888 | 0.0070433 | 0.0059546 | 0.0026351 | 44.2529
Lg1 | 0.01335 | 0.0044612 | 0.0032817 | 0.0025728 | 0.0036913 | 0.0041064 | 0.0032369 | 0.0040146 | 0.0027192 | 0.0033146 | 0.0047428 | 0.0036141 | 0.00071707 | 19.8407
Lg8 | 0.0001 | 0.0044612 | 0.0032817 | 0.0025728 | 0.0036913 | 0.0041064 | 0.0032369 | 0.0040146 | 0.0027192 | 0.0033146 | 0.0047428 | 0.0036141 | 0.00071707 | 19.8407
Lg1 + Lg8 | 0.01345 | 0.0089224 | 0.0065634 | 0.0051456 | 0.0073827 | 0.0082127 | 0.0064738 | 0.0080292 | 0.0054385 | 0.0066292 | 0.0094855 | 0.0072283 | 0.0014341 | 19.8407
Lg2 | 0.0006 | 0.0031066 | 0.0033245 | 0.0018105 | 0.0021033 | 0.0031257 | 0.00055618 | 0.0019911 | 0.0030709 | 0.0054864 | 0.0049277 | 0.0029503 | 0.0014614 | 49.5324
Lg3 | 0.0105 | 0.0030812 | 0.001329 | 0.0019981 | 0.0016517 | 0.0019747 | 0.0048811 | 0.0021527 | 0.0018213 | 0 | 0 | 0.001889 | 0.0014132 | 74.8122
Lg2 + Lg3 | 0.0111 | 0.0061878 | 0.0046535 | 0.0038086 | 0.0037551 | 0.0051003 | 0.0054373 | 0.0041438 | 0.0048922 | 0.0054864 | 0.0049277 | 0.0048393 | 0.00077672 | 16.0504
Fits whose peak areas vary, but whose sum is almost cosntant:
So we use the original ratios to diustribute the sum among the 2 peaks, giving:
Lb3 + Lb4 | Lb3 | Meas 1: Lb4; Meas 2: Sum
Original ratio | 0.63192 | Meas 1: 0.36808; Meas 2: 1
Resulting weights: | 0.055639 | Meas 1: 0.032408; Meas 2: 0.088047
Lb2 + Lb7 | Lb2 | Meas 1: Lb7; Meas 2: Sum
Original ratio | 0.99121 | Meas 1: 0.0087931; Meas 2: 1
Resulting weights: | 0.023914 | Meas 1: 5.0654e-05; Meas 2: 0.023965
Lg5 + Lg11 + L2N3 | Lg5 | Meas 1: Lg11; Meas 2: L2N3; Meas 3: Sum
Original ratio | 0.44 | Meas 1: 0.02; Meas 2: 0.54; Meas 3: 1
Resulting weights: | 0.00262 | Meas 1: 0.00011909; Meas 2: 0.0032155; Meas 3: 0.0059546
Lg1 + Lg8 | Lg1 | Meas 1: Lg8; Meas 2: Sum
Original ratio | 0.99257 | Meas 1: 0.0074349; Meas 2: 1
Resulting weights: | 0.0071745 | Meas 1: 5.3742e-05; Meas 2: 0.0072283
|  | Without fixing the peaks above: | [8 empty] | Fixing the peaks: