## Sheet1
Weights calibrated measuring a pure Ti standard, bulk, polished
Area of all peaks is fitted
A constant shift is maintained across all peaks
Sigma is fixed for all peaks
Areas of peaks La2, Lb17 and Lg5 are fixed on original weights because strongly overlapping with La1
Ni |  | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10
La1 |  | 2049.815 | 2080.617 | 2063.547 | 2054.565 | 2037.442 | 2054.453 | 2048.022 | 2056.75 | 2052.5 | 2048.427
Ll |  | 403.958 | 400.484 | 414.085 | 414.867 | 405.033 | 394.613 | 390.09 | 402.363 | 394.88 | 391.202
Ln |  | 96.467 | 102.867 | 91.634 | 97.305 | 102.666 | 107.298 | 109.142 | 104.078 | 106.867 | 117.705
Lb3 |  | 12.954 | 13.705 | 13.315 | 13.177 | 14.838 | 15.291 | 13.634 | 13.21 | 13.894 | 13.867
Lb4 |  | 12.954 | 13.705 | 13.315 | 13.178 | 14.838 | 15.288 | 13.621 | 13.211 | 13.894 | 13.867
Line | Original w | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5 | Meas 6 | Meas 7 | Meas 8 | Meas 9 | Meas 10 | Mean | Stdev | Er (%)
La1 | 1 | [10 empty] | 1 | bar: Meas 1 to Meas 10
Ll | 0.205 | 0.19707 | 0.19248 | 0.20067 | 0.20192 | 0.19879 | 0.19208 | 0.19047 | 0.19563 | 0.19239 | 0.19098 | 0.19525 | 0.0041755 | 2.1385
Ln | 0.0531 | 0.047061 | 0.049441 | 0.044406 | 0.04736 | 0.05039 | 0.052227 | 0.053291 | 0.050603 | 0.052067 | 0.057461 | 0.050431 | 0.0036758 | 7.2888
Ll + Ln | 0.2581 | 0.24413 | 0.24192 | 0.24507 | 0.24928 | 0.24918 | 0.2443 | 0.24376 | 0.24623 | 0.24446 | 0.24844 | 0.24568 | 0.0025195 | 1.0255
Lb3 | 0.0021 | 0.0063196 | 0.006587 | 0.0064525 | 0.0064135 | 0.0072827 | 0.0074429 | 0.0066572 | 0.0064228 | 0.0067693 | 0.0067696 | 0.0067117 | 0.00037699 | 5.617
Lb4 | 0.0128 | 0.0063196 | 0.006587 | 0.0064525 | 0.006414 | 0.0072827 | 0.0074414 | 0.0066508 | 0.0064232 | 0.0067693 | 0.0067696 | 0.006711 | 0.0003767 | 5.6132
Lb3 + Lb4 | 0.0149 | 0.012639 | 0.013174 | 0.012905 | 0.012828 | 0.014565 | 0.014884 | 0.013308 | 0.012846 | 0.013539 | 0.013539 | 0.013423 | 0.0007537 | 5.6151
The fitting of peaks Lb3 + Lb4 varies, but their sum is more or less constant |  | Meas 6: The fitting of peaks Ll + Ln varies, but their sum is more or less constant
So we use the original ratios to diustribute the sum among the 2 peaks, giving: |  | Meas 6: So we use the original ratios to diustribute the sum among the 2 peaks, giving:
| Lb3 | Meas 1: Lb4; Meas 2: Sum; Meas 7: Ll; Meas 8: Ln; Meas 9: Sum
Original ratio | 0.14094 | Meas 1: 0.85906; Meas 6: Original ratio; Meas 7: 0.79427; Meas 8: 0.20573
Resulting weights: | 0.0018918 | 0.011531 | 0.013423 |  |  |  | Resulting weights: | 0.19513 | 0.050545 | 0.24568
|  | Without fixing the peaks above: | [8 empty] | Fixing the peaks above: